Tribometer - Work Instruction

The following is a synthesis of the manuals and documentation provided by the manufacturer “TA Instruments”. Processes are amended to fit the laboratory environment at Tecnológico de Monterrey Campus Estado de México.

History of changes:

13Mar2019 - 1st version of the document

Edited by:

Antonio Osamu Katagiri Tanaka <A01212611@itesm.mx>

Revised by:

Nora Argelia Tafoya Medina <nuorriael@itesm.mx>

Dora Iliana Medina Medina <dora.medina@tec.mx>

Table of Contents

Startup Procedure	3

Shutdown Procedure	4

Install a Geometry	5

Install the Lower Stage	6

Remove the Lower Stage	8

# Start-up Procedure

- Ensure that the air supply is turned on.

- Turn on fluid circulation.

- Remove bearing clamp if fitted.

- Power on the system by pressing the power switch on the electronics box to the On “I” position.

- Connect the instrument to the software.

# Shutdown Procedure

- Turn off the power to the system by pressing the power switch on the electronics box to the Off “0” position.

- Turn off any fluid circulation.

- Fit the air bearing clamp.

- Turn off the air supply.

# Install a Geometry

- Ensure the compresses air supply is turned on.

- Ensure the instrument is powered on and initialized.

- Remove the bearing clamp by turning the draw rod counter clockwise

- Push the geometry up the spindle and hold.

Rotate the draw rod clockwise

Draw rod

Spindle

Geometry

The draw rod shall pull the geometry up into position. (tight but not forced)

# Install the Lower Stage

The following sections explain how to attach temperature modules using “Smart Swap”. The installation procedure is typically the same for all modules.

TIP: Make use of disposable gloves and tissue paper as some liquid may stain.
The lower stage attachments are stored in black containers/boxes.

Press the “release button” on the control panel, as described in the figures below.

Release button

A continuous green light indicates that the attachment can be fitted.
Notice that the release state will only stay active for 10 seconds after the “release button” is pressed.

Fit the attachment as shown below.

Ensure the mark and slot are align.

Mark

Slot

The installation shall fit as follows:

Connect the power cable and the fluid hoses.

Power cable

Fluid hoses

# Remove the Lower Stage

The following sections explain how to detach temperature modules using “Smart Swap”. The removal procedure is typically the same for all modules.

TIP: Make use of disposable gloves and tissue paper as some liquid may stain

Press the “release button” on the control panel, as described in the figures below.

Release button

Disconnect the power cable and the fluid hoses.

Power cable

Fluid hoses

- Press the “release button” again in the front panel.


Notice that the release state will only stay active for 10 seconds.

- Remove the attachment from the instrument.

- Store the removed attachment in its black container/box.